## Slide 1: Sensor Interfaces
Sensor Interface: directly connected to the sensor. Detects the typically small variations of the electrical quantity EX and convert it into a voltage, if required.
Further analog processing (amplification, filtering etc.)

## Slide 2: Sensors that produce a voltage
Output electrical quantity | Sensor type | Input physical or chemical quantity to be sensed.
Voltage | Thermoelectric sensors | Temperature difference Temperature Fluid flow rate Infrared radiation (bolometers) Gas concentration (catalytic sensors)
 | Electrochemical sensors | Ion concentration in electrolytes Gas concentration (e.g.. “lambda probes”)
 | Hall sensors | Magnetic Field Position (Proximity) Current measurement
 | Piezoelectric sensors | Force (ac detection) Acoustical pressure, acceleration.

## Slide 3: Sensors that produce a current or a charge
Output electrical quantity | Sensor type | Input physical or chemical quantity to be sensed.
Current | Optical sensors (photodiodes) | Infrared, visible and Ultraviolet radiation Imagers Proximity Opacity (e.g. smoke detectors)
Charge | CCD imagers | Visible radiation
 | High energy particle detectors | Ionizing radiation and particle detection

## Slide 4: Resistive sensors
Output electrical quantity | Sensor type | Input physical or chemical quantity to be sensed
Resistance | Thermistor and RTDs (Resistive Temperature Detectors) | Temperature Fluid flow rates Fluid velocity (e.g. hot wire anemometers) Gas concentration (catalytic sensors) Proximity
 | Piezo-resistors | Strain (strain gauges) Force (e.g. electronic scales) Pressure (barometers) Altitude Acceleration
 | Chemi-resistors | Gas or vapor concentration
 | Magneto-resistors | Magnetic field Proximity Orientation (e.g. electronic compass)
 | Photo resistors | Visible radiation

## Slide 5: Capacitive sensors
Output electrical quantity | Sensor type | Input physical or chemical quantity to be sensed.
Capacitance | Capacitive sensors (mechanical) | Acceleration Angular velocity (gyroscopes) Pressure
 | Capacitive sensors (chemical) | Gas concentration (e.g. humidity sensors)

## Slide 6: Frequently used sensor interfaces
Output quantity | Type of Interface | Notes
Voltage | Instrumentation Amplifier (In-Amp) | 
Resistance | Instrumentation Amplifier | Resistors should be mounted in a Wheatstone bridge configuration or biased by a current.
 | Trans-Impedance Amplifier (TIA) | Resistor must be biased with a voltage in order to produce a current
Current | Trans-Impedance Amplifier (TIA) | 
Capacitance | Trans-Impedance Amplifier (TIA) | Converting capacitance into a current by means of a periodic voltage waveform
 | Charge amplifier (switched capacitor) | 
Charge | Charge amplifier (switched capacitors) | 

## Slide 7: Instrumentation Amplifiers (In-Amps)
Required features:
Precise gain (AD)
High input resistance
Other important features
Differential input
Low input referred offset voltage
Low bias currents
Low input referred voltage and current noise
High CMRR (for differential amplifiers)
Large bandwidth
Vin+
Vin-
Vout
AD
Single-Ended
Instrumentation
Amplifier

## Slide 8: In-amps: errors due to the input noise/offset voltage and currents
Differential signal source
(Simplified Thevenin equivalent circuit)
Balanced source case (RS1=RS2=RS)
Useful Signal:

## Slide 9: Total voltage noise expressions
Balanced source case (RS1=RS2=RS)
DC components:
Noise components:
If in1 and in2 are uncorrelated and SI1= SI2 =SI:

## Slide 10: Discrete monolithic In-Amps
Typical pin configuration
One possible use of the SENSE terminal

## Slide 11: In-Amp with variable gain: Input and Output offset and noise
Typical two-stage architecture of In-amps
Generally, A2= 1, thus:
: input noise
: output noise

## Slide 12: Three-Op-Amp Instrumentation Amplifier
The circuit that provides VREF must have a very low output resistance (<<R)
Mismatch of these resistors degrades the CMRR of the second stage. Resistor trimming is necessary for CMRRs > 60 dB

## Slide 13: AD 620
At G≥100 the
amplifier BW is
larger than 100kHz.
The noise density
starts to fall for
f> 100kHz
The RTI noise
decreases with G

## Slide 14: AD 620
Nearly constant GBW product: BW determined by first stage
GBW does not increase much beyond the G=10 case: BW affected by second stage
Output noise >> input noise
Broad-Band Noise: √SBB
Low Frequency Noise
Integrated over 0.1-10 Hz
Broad-Band Noise: √SBB
Low Frequency Noise
Integrated over 0.1-10 Hz
Current
Slew -Rate
Settling times

## Slide 15
Input offset
Output offset
The effective input referred offset (RTI)
is a combination of the input and output offset
The input bias current and input offset current are similar, since
such a small bias current is the result of internal bias current cancellation
The AD 620 in-amp represents a good trade-off between input noise voltage, input bias currents and supply current (quiescent current)
AD 620

## Slide 16: Interfacing a sensor whose output is a current
(Norton equivalent circuit)
Goal:
to obtain a voltage proportional to the sensor current
IS is the quantity to be read, since it contain the useful information
ZS is the unavoidable output impedance of the sensor.
The ideal case is an infinite ZS (ideal current source)
Z is a property of the interface
and must be accurate.
In many cases Z is a resistance

## Slide 17: The simplest solution
iZ
Sensor
Interface
Simply connect a known impedance cross the sensor terminals

## Slide 18: Sensitivity vs Accuracy
iZ
Z should be large (magnitude) to obtain a high sensitivity
Z should be small (magnitude) to reduce the relative error (high accuracy)
For sensor marked by small values of the ZS impedance, finding a value of Z that satisfies both requirements is often impossible
Nominal law

## Slide 19: A better interface: the Trans-Impedance Amplifier (TIA)
iS
ZS
TIA
Vout
IZ
ZIN
Ideal case:

## Slide 20: The op-amp based TIA
operational
amplifier
0
iZ
iZ
Ideal case (perfect virtual gnd)
Sensitivity of the TIA stage
(transimpedance)

## Slide 21: TIA non-idealities
Finite gain
Finite input impedance
Typical dominant-pole frequency response
ZT

## Slide 22: TIA: input impedance ZIN
f0
Miler Effect:

## Slide 23: TIA: input impedance ZIN
f0 is defined as the frequency at which the amplifier gain magnitude is unity (0 dB)

## Slide 24: TIA input impedance

## Slide 25: Error due to the finite input impedance: (1) error on IZ
iZ
iZ
Ideal case:
Error on current IZ

## Slide 26: Error due to the finite input impedance: (2) error due to
iZ
Ideal case:
Relative
error contributions

## Slide 27: Reducing the 1/A error
iZ
Ideal case:
This voltage is affected only
by the error due to:

## Slide 28: TIA non-ideality: Noise and offset
Equivalent noise source of impedance Z
Thermal Noise:
Input referred noise voltage of the amplifier
Input referred noise current of the amplifier
For noise calculations, we will neglect finite gain effects (hypothesis of perfect virtual ground)

## Slide 29: Effect of in and vnR
in
in

## Slide 30: Effect of the voltage source

## Slide 31: Total output noise voltage and input referred noise current

## Slide 32: TIA used to read capacitive sensors
Differential capacitive sensor
The interface must read:
Balanced differential sensor
(e.g. MEMS accelerometer)
Pseudo differential sensor
(e.g. pressure sensor)
C0 is typically a large capacitance which does not vary with X, but is widely affected by process spread and often depends on temperature
physical quantity to be sensed

## Slide 33: Transformation of DC into a current
short-circuit current

## Slide 34: Use of the TIA to read current IS
Substituting the Norton equivalent circuit to the sensor, we have the usual TIA configuration, with Z=R
The useful signal DC(t) is modulated by sin(wSt): Demodulation is required to extract DC(t)
Synchronous demodulation allows detecting also the sign of DC(t)

## Slide 35: A few general considerations
Application of stimulus Vs
TIA
It is desirable to increase this factor to have a larger signal -to-noise ratio
VSM must be as large as possible (limited by Vdd)
wS cannot be too large because the error due to the finite impedance becomes not acceptable.
if wS increases:

## Slide 36: Complete interface
synchronous demodulator
gain = kD
(ideal TIA)

## Slide 37: Brief mention to TIA stability issues
Small phase
margin:
potential
instability
Introduces a zero that reduces the phase lag due to fb