## Slide 1
Magnetic Sensors
1) Hall sensors
Hall effect: Discovered by Edwin Hall in 1879
The Hall effect is based on the Lorentz force exerted on the moving charges (charge carriers) by a magnetic field B.

## Slide 2
Generation of the Hall voltage VH
Negative
charge
carriers
Positive
charge
carriers

## Slide 3
Integrated Hall Sensors
VH-
VH+
Ibias
SA=0.0412 V/T
for Ibias=0.5 mA
R0=2.2 kW

## Slide 4
Ex. 1: Electronic Compass
Earth’s Magnetic Field Intensity : 25-65 mT (0.25–0.65 G)
Horizontal component in Pisa (from map) 23300 nT ~ 23 mT

## Slide 5
Let us consider a maximum error of 5° :
BY
BX
Sensor arrangement:
Two orhogonal sensors that measure the X and Y components (horizontal components) of the magnetic field B
Angular resolution:

## Slide 6
Amplifier contribution to the angular error
AD 620
VOSI=15 mV VOSO= 200 mV (typical)
Offset current contribution: IioRS= 0.3 nA x 1.1 kW =0.33 mV << Vio-total
IBIAS=0.5 mA
SA=VH/B=0.0412 V/T
for B=23 mT : VH 0.95 mV
G=1000
OFFSET: affects the accuracy

## Slide 7
Noise: BW 0.1-100 Hz
Broad Band noise: eni=√SVBB=9 nV √ Hz  Integrated over BW=100Hz:
vnBB-rms=90 nV -> vnBB-pp=360 nV
Low frequency (Flicker) noise over 0.1-10 Hz -> 280 nV
Total noise voltage of the amplifier:
Offset referred to the measured magnetic field:
11 mT
The equivalent offset in the magnetic field measurement is much larger than the requirement (2 mT), Then the accuracy of this combination of sensor and in-amp is not adequate for an electronic compass
Let us see what we may expect in terms of resolution
Even considering only the contribution of the input noise voltage source, also the resolution is worse than required

## Slide 8
Amplifier contribution to the angular error (2)
Current noise: Broad-band component
Flicker component:
Noise voltage components due to the input bias current noise are
negligible with respect to the input noise voltage components.

## Slide 9
Sensor intrinsic offset
Bio25 mT
Bio(T)
DBio<5 mT (between 25 and 35 °C)
since 2 mT -> 5° angular error then:
Offset: -> 61° angular error
Drift: -> 12° angular error
Even considering only the offset
of the sensor, specs are not satisfied

## Slide 10
Typical readout approach for precise measurements:
Current Spinning
1
2
3
4
Ibias