## Slide 1: Sensor / object electrical coupling
The physical object is a solid, liquid or gaseous to which the quantity to be measured belongs
It can be a very simple object (e.g. a single-piece mechanical part) or a complex object, like the human body or a chemical reactor
The sensing operation requires the sensor / probe to interact with the physical object.
The sensor package (housing) must be designed to enable the required interaction and, at the same time, protect the sensors

## Slide 2: Sensor - object interaction
The interaction can be contactless (e.g. IR thermometer) or with contact (e.g. a thermocouple temperature probe).
Protection of the sensing elements is particularly important, for example, when the physical object is in liquid phase (liquid flow sensors, ion concentration sensors) or it is a corrosive gas.
Example: a liquid pressure sensor (10-bar resistant, usable for depth meters)
LPS33HW
Insulating gel:
Avoid direct contact with the liquid but ….
2) Transmit pressure from the liquid to the MEMS chip

## Slide 3: Electrical connection and continuity
A particular aspect that have to be assessed is the degree of electrical contact between the sensor terminals and the physical object
Only capacitive coupling
(AC coupling)
Only resistive coupling
C: order of pF
VISO (maximum that the barrier can withstand)
The presence of a dc path can be unwanted (defective insulation) or necessary

## Slide 4: Examples
H2O
conductivity probe
terminals
Probe for the measurement of the water conductivity
(used to monitor the quality of water supplies)
In this case, electrical continuity is necessary to perform the measurement
ceramic or plastic
insulator
thermocouple
(or thermistor)
metal case
liquid
Electrical continuity is not necessary, but a leakage conductance between the sensor terminal and the liquid can develop when the insultaor properties degrade with time
Temperature probe

## Slide 5: Voltages due to mains inside a building
power wall socket
N= Neutral
E= Earth
L= Live (or "phase")
The neutral is connected to the earth at the distribution transformer

## Slide 6: Consequences of the presence of AC power distribution
Generally:
B1: conducting body

## Slide 7: Particular case: the power supply
Isolation barrier
Power
In any power supply it is impossible to avoid leakage impedances (ZGL, ZGN) between the input output terminals
Vdd
gnd
ZGF
VGF
This is the actual equivalent circuit that takes into account also the connection towards the floor and the relative voltage.

## Slide 8: A few significant cases
Presence of electrical continuity between a conductive physical object and the sensor terminals.
Suppose that: 1) the input impedance is infinite
2) the voltages between the electrodes and the liquid are negligible
With this simplified model, the input common mode voltage of the in-amp is:
Since VBF and VGF are important fraction of the mains voltage (220 V r.m.s):
The amplifier will be damaged or the protection at the input terminals activate, making the measurement meaningless
Example

## Slide 9: Two possible solutions:
IG
The problem with solution is that a current flows trough one of electrodes. Since the series resistance between the electrode and the solution can be relatively high, the current can cause an important error in the measurement.
Connect one of the two terminals to gnd. In this way, the input common mode voltage is close to zero.
connection of one input to gnd

## Slide 10: Second solution
Use a third electrode to connect the gnd of the amplifier to the fluid
("Three electrode measurement")
IG
IG
VB
VB

## Slide 11: Second solution
If Z3 << ZBF+ZGF, the effect of voltages VBF and VGF on VC can be as small as a few hundred mv
If the input amplifier impedance is >> Z1,Z2:
The differential voltage is not affected by VBF and VGF
VB
VB
Such a small VC can be tolerated by the amplifier

## Slide 12: Effect of the finite input impedances of the amplifier
(this is an AC
disturbance voltage)
Due to unavoidable impedance mismatch, the two impedance ratios a1 and a2 are not equal and a common mode disturbance voltages generate a differential voltage.
a1
a2

## Slide 13: Effect of the finite input impedances of the amplifier
VB can include the fundamental component of the AC power line (50 or 60 Hz, depending on country) but also harmonics at higher frequency and other disturbance signals from motors, DC-DC converters, etc,
The amplifier input impedances Zi1 and Zi2 are mainly capacitive, while Z1 and Z2 may include both capacitive and resistive components.
Clearly, in the case that Zi1,Zi2 >> Z1,Z2, the impedance ratios a1 and a2 are very close to 1 and their difference is close to zero. This is the case of infinite input amplifier that we have analyzed before, where the effect of VB on Vid was null.

## Slide 14: General case of no electrical continuity
Here, we represent a typical connection between amplifier and sensor ...
... and we add a disturbance voltage applied between two conductors. This disturbance can be the AC power line, but also have a different origin
Circuit for calculation of the disturbance at the amplifier input
That can be further simplified, applying the Thevenin theorem at the bridge diagonal where Zp-in is placed. The Thevenin equivalent does not include Zp-in
These impedances are usually capacitances

## Slide 15: General case of no electrical continuity
VTH is zero only for:
which generally does not occur.
With a careful design:
For a given Zp-in (sensor and amplifier), the higher Zth, the smaller the input disturbance

## Slide 16: Disturbance minimization: capacitive (electrostatic) coupling
The magnitude of this component, due to the "hot" terminal, is much higher than the other, because capacitances that involve the gnd node (and all other nodes connected to it) are much larger.
In order to minimize the interference, we can maximize the impedance between the "hot terminal" and the disturbance conductors. This means reducing the related capacitances CS1 and CS2.

## Slide 17: Mutual capacitance reduction: shielding
We add a conductive shield that wraps around the sensitive input node (and other sensitive nodes, if any)
External conductors cannot produce fields inside the shield ("faraday cage")
For relatively long connections between the sensor and the amplifier, a shielded coaxial cable can be used to protect the internal "hot" terminal from interaction with disturbance conductors.

## Slide 18: Example
Oscilloscope probe. The input "hot" terminal is shielded.
Just adding 6 cm of unshielded conductor ...
~10 mV p-p
50 Hz with strong harmonics
The exposed conductor is less than 1 cm long

## Slide 19: Spectral content of the disturbance voltage
50 Hz
horizontal scale: kHz
The voltage picked up from the AC power line by capacitive coupling is much more distorted than the original voltage, since high order harmonics find much less attenuation:
At low frequencies the resistive component dominates

## Slide 20: Advantage and limit of differential reading.
With a symmetrical routing of Vip and Vin, we can obtain: CP1=CN1, CP2=CN2
In a differential connection, the signal is the voltage difference between two wires, none of which is gnd.
The sensor has three terminals.

## Slide 21: Advantage and limit of differential reading.
equal
equal
Nd1
Thanks to the symmetry:
Then, to obtain Vid=Vip-Vin = 0 we need a balanced source and balanced amplifier:
There are critical cases where it is not possible to have a balanced source.
A pseudo-differential configuration is not effective:
Nd1
We apply three times the Thevenin equivalent between Nd1 and:
Vip, gnd, Vin, respectively.

## Slide 22: Interference from variable magnetic fields
The loop formed by the mesh that includes the sensor output port and amplifier input port is subjected by an electromotive force given by:
The current in the loop is given by:
For a sinusoidal magnetic field:
The voltage at the amplifier input is then:
If |ZiA|>>|ZoS|
In order to reduce the disturbance, it is necessary to reduce the area of all critical loops. Twisting the cables reduces the equivalent area of connection cables. Note that magnetic shield are much less effcetive than electrical ones.

## Slide 23: Ground loops
The picture below depicts two reaout channels (two sensors). The ground of the two channels must be connected together in a single point. Connecting the grounds in two spatially separated points, creates a loop that can sustain considerably large currents.
Ground connection 1
Ground connection 2
Zloop is the series impedance of the whole loop.
Since ZG can be a large fraction of Zloop, an important portion of the electromotive force falls across ZG. The voltage VG gets added to the output signal of the sensor.