## Slide 1: Hazards due to unwanted coupling with high voltage sources
Hazards
For the sensor system (risk of permanent damage)
Risk for humans (injuries or death)

## Slide 2: Premise: resistive current sensing in power systems
For negligible
power loss:
Vsense  mV
Amplification
is needed
Input VCMVsup
Very high CMRR is required
if Vsup has large variations
High-side current sensing
Low-side current sensing
+) Amplifier input VCM0
-) Load gnd is not power supply gnd
-) Current sensing of independent loads cause gnd voltage differences
different gnds

## Slide 3: Risk for the sensor system
VH=NVB
Rsense
VB
Conversion
Processing
Display ....
Imotor
Icharge
3.7 V
 100-800 V
VV
VI
Goal: monitor the current through the battery and the voltage of individual cells for State-of-Charge estimation.
Voltage and current sensing in battery packs for e-vehicles
Such a high common mode voltage at the input damage the amplifiers!
High-side current sensing

## Slide 4: Hazards to humans
Case 1: AC powered biomedical apparatus, electrically connected to the patient.
Let us suppose that this impedance gets much smaller due to a failure in the transformer insulation.
Depending on ZBF, a large current may pass through the body. A current that passes through the heart is potentially fatal.

## Slide 5
Hazards to humans
Case 2: Sensing systems with possible continuity with a high voltage domain.
The sensor is isolated from conductors of the high voltage domain, but a leakage impedance may originate from dielectric failure or an unwanted contact (Zleak1).
From the operator side, the front panel control may lose insulation due to wearing or presence of moisture / accidental liquid pouring

## Slide 6: Isolation approach 1: Isolation Amplifier
Double protection concept: this is a strict requisite in biomedical instrumentation
Probe (electrode)
or sensor
Isolation amplifier

## Slide 7: Techniques for analog signal insulation
Transformer coupling
Optical Coupling
Capacitive Coupling
Example of transformer coupling: the AD210

## Slide 8: Principle of analog optical coupling
Example: Broadcom ACNT-H79A
An optical link is marked by strong non-linearity. It can be used in an open-loop scheme only to transmit digital signals.
For analog signals, a feedback scheme can be used, as in the figure. The LED CR1 is coupled in a symmetrical way to photodetectors CR3 and CR2, thus:
virtual short circuit due to optical feedback at A1.
(photocurrents)

## Slide 9: Isolation amplifier with capacitive coupling
Data that cross the barrier are digital (bitstream). Differential data links are used
Data are converted back to analog by filtering.

## Slide 10: Isolation approach 2: Isolated Digital Link
Digital Data Isolator
Advantages of digital isolation
Does not affect accuracy
Cost is generally lower than analog
Size is also generally smaller
Digital isolators are based on two principles:
(1) Optical (photocouplers) (2) Capacitive

## Slide 11: Example of digital data isolator: I2C isolator
Digital data are modulated by a high frequency carrier (on-off modulation (on-off keying, OOK) and transmitted across the barrier through a capacitive barrier.
Texas Instrument
4 channels are necessary because each one of the two I2C lines is bidirectional

## Slide 12: Example of isolated power supply
RECOM REM10 (10 W)
Medical compliant
For standby mode
Simbolo
Isolation voltage: 5 kV