## Slide 1: Importance of detecting biopotentials
Diagnostic purposes: Detecting a possible disease (e.g. heart and neural pathologies, etc.)
Implantable, wearable biomedical devices:
Cardiac pacemakers (on demand stimulation)
Control of limb prostheses (smart prostheses)
Advanced human-machine interfaces
Early detection of pathological situations (e.g. , automatic defibrillator, migraine attacks) for on-demand drug delivery or electrical stimulation)
Research: understanding the functioning of human or animal body, with particular interest in functioning of the brain

## Slide 2: Type of biopotentials and related challenges
ECG: Electrocardiogram, produced by the electrical activity of the cardiac muscle
EEG: Electroencephalogram, produced by the neurons in the brain
EMG: Electromyogram, produced by the contraction and relaxation of muscles
ERG: Electroretinogram, produced by the activity of the retina.
Main biopotentials, according to the source

## Slide 3: Biopotential acquisition system
Electrodes
AFE
Isolation
Processing
Number and position of the electrodes
Type of electrodes
Differential input
High CMRR
Specifications of:
Bandwidth
Noise
Input bias currents
Artefact correction capability
Impedance measurement (quality of contact)

## Slide 4: Detection of biopotential through non-invasive electrodes
Principle: electrodes are placed on the skin
skin
fiber
currents
Electrically excited nerve or muscle fibers induce currents in the tissues that produce potential differences across different points of the body
electrodes
To the electronic interface

## Slide 5: Electrodes for biopotential detection: the Ag / AgCl electrode
Plastic cup
Silver (Ag)
Silver chloride (AgCl)
Electrolyte gel
adhesive layer
out
Skin
The Ag/AgCl electrode is one of the best option, Can be disposable or reusable.
Tissues
Skin / tissues
(ionic currents)
Terminal
(electron currents)
Vhc is the half-cell potential. It introduces a dc offset that can be critical for the AFE
Electrode equivalent circuit

## Slide 6: Other types of electrodes
Metal electrodes:
Can be advantageous when the electrode must be reused, or the electrode position must be changed frequently. Need application of electrolyte gel.
Example: cuff electrode
rubber cuff
metal
electrode
Dry electrodes
Do not need application of electrolyte gel.
Capacitive electrodes
Microneedle arrays
Conductive silicone electrodes
Microneedles
Insulator (touchless) or conductive silicone rubber
Active electrode

## Slide 7: Electrodes for in-depth biopotential monitoring
Sometimes, it is necessary to detect the biopotential in a position very close to the organ that is generating it. In this case, invasive catheter or needle electrodes are required. They do not need gel, since internal tissues are wet
Examples:
Needle electrode
(used to contact muscle fibers directly)
conductor
insulator
Catheter: used to pick up potential differences within organs (e.g. the heart).
Can be brought in place through veins or arteries
insulator (flexible)
Metal electrodes

## Slide 8: Electrode position and number: ECG
C
LA
RA
RL
LL
C; chest (typically 6 positions, V1-V6)
RL: Right Leg
LL: Left Leg
RA: Right Arm
LA: Right Arm
Reading is differential with the right-leg always used as gnd (necessary to set the input common mode voltage)
This is an example of connection
("Lead I")

## Slide 9: ECG example of leads
Lead I : LA-RA
Lead II : LL-RA
Lead III : LL-LA
Augmented:
aVR
Chest V4
Different leads give different diagnostic information.
For monitoring purposes, a single lead is generally used
Resistors (placed inside the ECG instrument, are used to average the voltage of two or three electrodes

## Slide 10: EEG and EMG electrode placement
Each electrode position has a code. Electrodes in the "middle sagittal plane" are generally used as gnd.
EEG
EMG
Surface EMG:
Compact bipolar
electrode for EMG
middle sagittal plane
10-20 electrode system: refers to the spacing between electrodes in terms of percentage, 10 % 20 %
(8 - 128 electrodes)

## Slide 11: Biopotentials characteristics
Type of biopotential | Bandwidth | Typical level
ECG | 0.05 Hz - 100 Hz | 1 mV -5 mV
EEG | 0.5 Hz - 40 Hz | 10 mV - 100 mV
EMG | 20 Hz -2 kHz | 1 mV - 10 mV
ERG | 1 Hz - 1 kHz | 10 mV - 100 mV
Very low frequency:
high pass filter is critical and exhibits long settling times (due to poles at very low frequency)
Very small magnitude. Requires very-low noise amplifiers and careful rejection of interferences
Range includes 50/60 Hz: prone to interference from power line

## Slide 12: Power line interference
Electrode
impedance
Z1
Z2
Z3
Due to mismatch of electrode impedance (Z1,Z2), a1≠a2 and VB produces a differential input
Z3 may be large enough to produce a VB of hundred mV
iG

## Slide 13: Right Leg Drive (RLD) Method
The common mode voltage of the in-amp input terminals is extracted and amplified with a high-gain inverting amplifier (ARL)
Thanks to the high loop gain, virtual ground is established at the input of ARL, meaning that Vic is strongly attenuated.
Extraction of the common mode voltage from an in-amp

## Slide 14: Right Leg Drive (RLD) Method
Then, the strong reduction of the input common mode voltage results in a corresponding reduction of the differential mode disturbance voltage caused by the common mode.
am

## Slide 15: Electrode offset
Tissue
Electrodes
Due to mismatch in the electrodes, their degradation with time, or to local differences in the electrolyte composition, the half-cell potentials can be different. DC voltage up to several tens mV may appear

## Slide 16
Electrode offset
The DC offset slowly varies with time, producing the so-called "baseline-drift"
This phenomenon can be very important, distorting the useful signal and causing saturation of the amplifier.
The DC component must be removed with a high pass filter.
ECG | 0.05 Hz - 100 Hz
For example, ECG recording would require a filter with a lower cut-off frequency og 0.05 Hz

## Slide 17: Versatile AFE schematic
Three electrode scheme
For fL=0.05 Hz, t=3.18 s 5 t  16 s
Such a long settling time can lead to long period of unavailability
Programmable Gain Amplifier
Programmable Low Pass Filter

## Slide 18: Baseline fast recovery circuit
Without recovery circuit
Amplifier saturation caused by a disturbance
Switch S1 is activated for a fixed time
S1

## Slide 19: Bibliography
[1] Khan University, "Signal propagation: The movement of signals between neurons", https://www.khanacademy.org/test-prep/mcat/organ-systems/neural-synapses/a/signal-propagation-the-movement-of-signals-between-neurons
[2] University of Texas, Neuroscience Online, https://nba.uth.tmc.edu/neuroscience/m/s1/index.htm
[2] Lumen, Anatomy and Psychology: Neuromuscular Junctions and Muscle Contractions, https://courses.lumenlearning.com/cuny-csi-ap-1/chapter/neuromuscular-junctions-and-muscle-contractions/
[3] Analog Devices, " Biopotential Electrode Sensors in ECG/EEG/EMG Systems" https://www.analog.com/en/technical-articles/biopotential-electrode-sensors-ecg-eeg-emg.html#:~:text=A%20biopotential%20electrode%20is%20a,ion%20current%20to%20electron%20current.